## Technical Requirements
Technical Requirements of Medical Device Product
Prepared for: |  | CIRS | Prepared by: | CIRS_Edwin |  | Updated data: 9th June 2017
Performance |  |  |  | Test Methods
No. | Sub_Code | Index | Criteria/Threshold | No. | Sub_Code | Test Methods
1 | Product Model & Specifications
| 1.1 | Product Model
| 1.2 | Product Specifications and the Division Instruction
| 1.3 | Product Structure Diagram
2 | Performance |  |  | 3 | Test Methods
| 2.1 | Appearance | Colorless , transparent viscous gel, without any visible foreign matter (HA Filler) |  | 3.1 | Cross-linked sodium hyaluronate gel is placed vertically, under the illumination of 1000 lx, and then rotates it randomly; observe the gel from the horizontal direction.
| 2.2 | valid usage amount | should be between 90% - 120% of the labeled amount |  | 3.2 | Take out the cross-linked sodium hyaluronate gel by normal use method, weigh it. Then divide the weight of the gel by the density of gel (ρ=1.01 g/ml).
| 2.3 | Particle size distribution | gel particles distributed between 100um and 1000um shall occupy more than 90% |  | 3.3 | Conduct test according to the standard or laser particles distribution instrument
| 2.4 | injection force | the maximum force shall be less than 40 N and the average shall be 13±5 N |  | 3.4 | Conduct the injection force test according to the standards
| 2.5 | elastic modulus G' viscous modulus G'' | Under 1 Hz of test frequency, the elastic modulus G' shall be 400±100 Pa, and the viscous modulus G" shall be 900±150 Pa |  | 3.5 | Using Malvern rheometer and operate under (25 ± 0.2) ºC, at a shear rate of frequency sweep from 0. 1 Hz-10Hz, plot the logarithmic coordinate graph of viscous modulus G "and the elastic modulus G 'in terms of logarithm and frequency, record the value at the vibration frequency of 1 Hz
| 2.6 | infrared identification | there shall be characteristic infrared absorption peaks near the 617±5, 1035±5, 1082±5, 1154±5, 1409±5, 1638±20, 2895±5, and 3434±200 of FT-IR (cm-1) |  | 3.6 | Conduct infra-red identification test according to standard
| 2.7 | swelling degree | the swelling degree of cross-linked sodium hyaluronate gel shall be more than 30 |  | 3.7 | Conduct swelling degree test according to standard
| 2.8 | osmotic pressure | Should be between 270 mOsmol / L - 350 mOsmol / L |  | 3.8 | Direct sampling and conduct test according to the provision of appendix IXG of Chinese Pharmacopoeia (2010)
| 2.9 | PH value | Should be within the range of 6.8-7.6 |  | 3.9 | Directly use sodium hyaluronate gel according to the provision of appendix VIH of Chinese Pharmacopoeia (2010)
| 2.10 | Content | the content of sodium hyaluronate shall be 90%-120% of labeled amount |  | 3.10 | Conduct the test of content according to standard
| 2.11 | protein content | Protein content should be not more than 0.1 % |  | 3.11 | Conduct the test of protein content according to standard
| 2.12 | intrinsic viscosity | the value shall be more than 100 cm3/g |  | 3.12 | Conduct the test of intrinsic viscosity according to standard
| 2.13 | total amount of heavy metal | (calculate Pb) no more than 5 μg / g |  | 3.13 | Measure the content of total heavy metal (calculate by Pb) according to the provision of appendix VIIIH of Chinese Pharmacopoeia (2010)
| 2.14 | cross-linked agent residual (BDDE) | Residual amount of the crosslinking agent is not more than 2.0μg / g |  | 3.14 | Conduct the test of BDDE residual according to standard
| 2.15 | free sodium hyaluronate content | the value shall be less than 20 mg/ml |  | 3.15 | Conduct the test of free sodium hyaluronate content according to standard
| 2.16 | EtOH residual | the value shall be not more than 400 μg/g |  | 3.16 | Conduct the test of EtOH residual according to standard
| 2.17 | sterility test | The product should be sterile |  | 3.17 | Cross-lined sodium hyaluronate gel: measure it according to appendix XIH of Chinese Pharmacopoeia (2010)
| 2.18 | bacterial endotoxin | There should be ≤0.5EU / ml |  | 3.18 | Bacterial endotoxin: measure it according to appendix XIE of Chinese Pharmacopoeia (2010)
| 2.19 | hemolytic streptococcolysin (fermentation method) | There should be no hemolysis |  | 3.19 | Biological fermentation: put 1ml of medical sodium hyaluronate gel on blood agar plate, cultivate it for 24 h, 37±1ºC
| 2.20 | Syringe , needle | the injection syringe should accord with requirement of GB 15810 |  | 3.20 | Test it according to GB.15810-200